# Design Document

## Raspberry Pi Camera Mount for Slit Lamp Observer Tube: Measurement, Design, and 3D Printing Workflow

### Purpose of This Document

This document describes the complete workflow for designing a mechanical mount to attach a Raspberry Pi camera module to the observer (monocular) tube of a slit lamp microscope. It covers required measurements, design constraints, mechanical concepts, CAD modelling steps, and preparation for 3D printing or commercial manufacturing.

The intent is to enable repeatable design, testing, and iteration of a camera mount suitable for proof-of-concept and demonstration use.

### Design Philosophy

The mount must satisfy the following high-level principles:

- The mount must be non-destructive and fully reversible.

- The mount must not apply significant mechanical stress to the observer tube.

- The camera sensor must be precisely aligned with the eyepiece exit pupil.

- The design must tolerate small dimensional variations between slit lamps.

- The design must support iterative adjustment during optical alignment.

The mount is not a structural component of the slit lamp and must not interfere with normal clinical operation.

### System Overview

The mounting system consists of four functional elements:

- An eyepiece interface that mates with the slit lamp observer eyepiece.

- A camera holder that securely positions the Raspberry Pi camera.

- An axial spacing mechanism that positions the camera sensor at the exit pupil.

- A locking or retention mechanism that maintains alignment once set.

These elements may be integrated into a single printed part or split into multiple components depending on design preference.

### Initial Measurements Required

Before any CAD design begins, the following measurements must be taken on the actual slit lamp. These measurements should be taken using digital calipers where possible.

- Eyepiece and Observer Tube Measurements

- Outer diameter of the rubber eyecup.

- Outer diameter of the eyepiece barrel beneath the eyecup.

- Length of the eyepiece exposed above the observer tube.

- Whether the eyepiece can be removed from the observer tube.

- Angle of the observer tube relative to horizontal.

Optical Position Measurements

- Approximate distance from the top of the eyepiece to the comfortable eye position of an observer.

- Observed vignetting behaviour when moving the eye forward and backward.

Camera Module Measurements

- Camera PCB width and height.

- Distance from camera PCB face to camera sensor plane (from datasheet or measurement).

- Overall thickness of the camera module including connector.

- Ribbon cable exit direction and bend radius.

All measurements should be recorded with tolerances where possible.

### Derived Design Targets

From the measurements and optical principles, the following targets should be defined:

- Target distance from eyepiece top surface to camera sensor (typically 10 to 15 millimetres).

- Acceptable axial adjustment range (minimum plus or minus 3 millimetres).

- Acceptable lateral alignment tolerance (approximately plus or minus 0.5 millimetres).

- Minimum clear aperture through the mount (at least 8 to 10 millimetres).

These targets define the internal geometry of the mount.

### Eyepiece Interface Design

The eyepiece interface is the most critical mechanical feature.

Recommended approach:

- A cylindrical sleeve that fits over the eyepiece rubber cup or barrel.

- Internal diameter sized for a friction fit, not a clamp.

- Wall thickness sufficient to resist deformation but flexible enough to slide on.

Design considerations:

- Allow for rubber eyecup compression.

- Avoid sharp internal edges that could damage the eyepiece.

- Prefer a smooth internal bore.

Optional enhancements:

- Internal groove for a thin rubber ring to improve grip.

- Split-sleeve design to allow slight expansion.

### Camera Holder Design

The camera holder must position the camera sensor precisely on the eyepiece optical axis.

Recommended approach:

- Flat mounting face for the camera PCB.

- Defined reference plane corresponding to the sensor position.

- Mechanical features to prevent rotation of the camera module.

Design considerations:

- Account for ribbon cable routing and strain relief.

- Ensure the camera lens or aperture is centred on the mount axis.

- Avoid obscuring the camera field with the mount geometry.

Axial Adjustment Mechanism

Axial adjustment is required to fine-tune exit pupil alignment.

Recommended approaches:

- Threaded barrel design where the camera holder screws in and out.

- Spacer ring system with multiple thin shims.

- Sliding tube with locking set screw.

For proof-of-concept work, a threaded design is recommended because it allows continuous adjustment without reprinting parts.

Thread design notes:

- Use coarse threads to reduce print sensitivity.

- Allow sufficient thread engagement to prevent wobble.

- Design a hard stop to prevent over-insertion.

### Alignment and Locking

Once correct alignment is achieved, the system must be locked in place.

Recommended locking methods:

- Thread friction plus hand-tightening.

- Small nylon set screw pressing against a flat surface.

- External locking collar.

Avoid metal fasteners contacting optical components.

### CAD Modelling Workflow

Recommended CAD approach:

- Model in parametric CAD software (e.g. Fusion 360, FreeCAD, OpenSCAD).

- Define key dimensions as parameters (eyepiece diameter, axial offset, camera offset).

- Model components in the following order;

- Eyepiece sleeve

- Axial adjustment mechanism

- Camera holder

- Test-fit virtually by assembling components in CAD.

- Export models as STL files for printing.

### Material Selection for 3D Printing

Recommended materials for initial prototypes:

- PLA for fast iteration and low cost.

- PETG for improved toughness and temperature resistance.

Material considerations:

- PLA is adequate for indoor, short-term use.

- PETG is preferred for repeated handling and hospital environments.

Avoid brittle resins or very flexible materials for early iterations.

### Print Settings and Tolerances

Recommended print settings:

- Layer height: 0.2 millimetres.

- Wall thickness: minimum 1.2 millimetres.

- Infill: 20 to 30 percent.

Design tolerances:

Add clearance of 0.2 to 0.4 millimetres on friction-fit diameters.

Test-fit and iterate rather than relying on nominal dimensions.

13.	Iterative Testing Process

The mount design should be validated iteratively:

Print initial version with generous clearances.

Mount camera and perform optical alignment.

Record vignetting and centring performance.

Adjust axial position and note optimal sensor distance.

Update CAD model to refine dimensions.

Reprint and retest.

Multiple iterations are expected and normal.

14.	Preparation for Commercial Printing

Once a stable design is achieved:

Confirm all dimensions against final measurements.

Export final STL or STEP files.

Specify material and colour preferences.

Provide orientation guidance to the printer if required.

Commercial printing is suitable once design iteration has stabilised.

15.	Summary

Designing a Raspberry Pi camera mount for a slit lamp observer tube is primarily a mechanical alignment challenge. By carefully measuring the eyepiece geometry, designing for axial adjustment, and iterating through test prints, a stable and repeatable mount can be achieved without modifying the slit lamp.

This workflow supports both in-house 3D printing and outsourced commercial manufacturing.

⸻

If you want next, I can:

•	Turn this into a step-by-step checklist you can literally follow in the workshop

•	Provide example parametric dimensions you can drop straight into CAD

•	Generate a first-pass OpenSCAD model based on assumed dimensions

•	Adapt this for a specific slit lamp brand or Raspberry Pi camera version

Just tell me.