January 18, 2026

Fabian Brandimarte

FDB Projects

Project Outline and Execution Plan (Proof of Concept): Slip Lamp Observer View Recording System

### Executive Summary

This document describes a phased proof-of-concept project to explore the feasibility of a low-cost, removable camera system capable of recording video through the observer monocular of an ophthalmic slit lamp. The intent is to capture exactly what the clinician sees during examination, enabling video documentation, teaching, and specialist communication using readily available components.

The project deliberately focuses on technical feasibility, usability, and image adequacy, and is not intended for diagnostic use, routine clinical deployment, or regulatory submission. No patient-identifying data is stored or processed as part of this proof-of-concept.

To minimise risk and avoid premature optimisation, the work is structured into clearly defined phases, each addressing a specific uncertainty:

•	reliable headless hardware control via a single illuminated button

•	optical feasibility of observer-view imaging

•	mechanical stability and repeatable alignment through a removable mount

•	clinician workflow integration and data handling

Each phase includes explicit decision gates and off-ramps, allowing the project to be paused or concluded early if key assumptions are not validated. This ensures that all work undertaken yields standalone technical insight and reusable learning, even if later phases are not pursued.

Mechanical design and 3D-printed mounting solutions are intentionally deferred until optical feasibility has been demonstrated using temporary, fully reversible fixtures. Where required, parametric CAD modelling and rapid prototyping are planned using open-source tools and consumer-grade additive manufacturing to enable fast iteration and low cost.

The outcome of this proof-of-concept is expected to be a demonstrable, non-diagnostic recording system and a clear evidence base to inform whether further development, refinement, or exploration is justified.

### Project Overview

Problem Statement: “Current slit-lamp imaging solutions force clinicians to choose between cheap but awkward smartphone adapters and expensive, closed OEM camera systems. There is no affordable, dedicated, workflow-friendly way to record exactly what the clinician sees, especially as video.”

Purpose: The purpose of this project is to design and validate a low-cost, removable camera system that records video through the slit lamp observer monocular, capturing exactly what the ophthalmologist sees during examination, enabling sharing of the video and still images with other specialists.

Objective Framework: This project is structured as a phased proof-of-concept, where each phase de-risks a specific technical, optical, or usability aspect before progressing further. Objectives are defined at both a programme level and a phase level, allowing incremental validation without premature optimisation or regulatory overreach.

Primary Objective (Programme-Level): To demonstrate the technical feasibility, usability, and image adequacy of a low-cost slit-lamp observer-view video recording system using readily available components, prior to any clinical certification or productisation.

Phase Level Objectives

- Objective 1: Electronics & Control (Hardware Proof of Concept)

- To demonstrate reliable, deterministic physical control of a headless camera system using a single illuminated button, including robust start/stop behaviour and clear recording-state feedback, implemented on a Raspberry Pi–based platform.

- Objective 2: Optical & Imaging Feasibility

- To demonstrate that usable slit-lamp observer-view video can be captured through the monocular eyepiece using a compact CS-mount camera, without significant vignetting, distortion, or loss of clinically relevant detail.

- Objective 3: Workflow & Usability

- To validate that the system integrates naturally into an ophthalmologist’s slit-lamp workflow, requiring minimal cognitive load, no interaction with screens or software during examination, and no disruption to standard clinical practice.

- Objective 4 – Data Handling & Demonstration Readiness:

- To demonstrate predictable and safe handling of recorded video files, including automatic storage to removable media and straightforward retrieval for non-diagnostic review, sharing, or teaching.

Scope: This project is strictly a proof-of-concept. It is not intended for diagnostic use or routine clinical deployment without further regulatory, safety, and data protection assessment.

### Success Criteria

The proof-of-concept will be considered successful if all the following conditions are met:

The system is capable of recording clear and stable video through the slit lamp observer eyepiece.

Recording can be started and stopped using a single physical button.

Video files are automatically saved to removable USB storage.

The system operates headless, without keyboard, mouse, or monitor.

The mounting solution does not damage the slit lamp, is removable, and maintains repeatable optical alignment.

### System Architecture

Hardware Components:

- Raspberry Pi Zero 2 W

- Camera module (OV5647 or Raspberry Pi Camera v3)

- Physical push button connected via GPIO

- Status LED for recording state feedback

- USB storage device

- Power supply (USB wall adapter or battery pack)

Software Components:

- Raspberry Pi OS Lite

Python-based control service running at boot

libcamera for video capture

GPIO interrupt handling

Automatic file naming and storage management

### Overview of project phases and design

The work described in this document is structured as a series of clearly defined, sequential phases, each designed to validate a specific assumption or de-risk a particular aspect of the system before further effort is committed.

The phases are intentionally ordered to progress from low-cost, low-risk validation toward more involved mechanical and workflow development. Advancement between phases is contingent on meeting defined acceptance criteria.

This phased approach ensures that:

- technical complexity is introduced only when justified by results

- effort scales gradually rather than prematurely

- failure at any stage produces useful, standalone learning

Each phase addresses a distinct question:

| Phase | Question |
| --- | --- |
| Phase 1 – Hardware Bring-Up | Does the core hardware function reliably in a headless configuration? |
| Phase 2 – Button-Controlled Recording Logic | Can the system be controlled deterministically and intuitively using a single physical interface? |
| Phase 3 – Slit-Lamp Integration (Optical and Mechanical) | Can observer-view imaging be achieved and maintained through an integrated, removable housing that combines optical alignment and mechanical stability using low-cost components? |
| Phase 4 – File Handling and Data Flow | Can recorded data be stored, retrieved, and shared predictably and safely? |
| Phase 5 – Reliability and Workflow Enhancements | Is the system sufficiently robust and intuitive for demonstration in a clinical environment? |

### Phase 1 – Hardware Bring-Up

Objective: To verify that all hardware components function correctly in a bench-top setup.

Tasks:

- Install Raspberry Pi OS Lite.

Enable the camera interface.

Verify camera operation using libcamera.

Wire and test the GPIO button input.

Verify that USB storage is automatically mounted.

Deliverables:

- The Raspberry Pi boots reliably in headless mode.

The camera preview works consistently.

Button presses are detected correctly in software.

USB storage is recognised without manual intervention.

Acceptance Test: A 10-second test video is successfully recorded and saved to USB storage.

### Phase 2 – Button-Controlled Recording Logic

Objective: To enable simple, clinician-friendly control using a single physical button.

Button Behaviour:

- A short press starts recording.

- A short press while recording stops recording.

- A long press of five seconds or more triggers a safe system shutdown.

Deliverables:

- A Python service runs automatically on boot.

Recording starts and stops reliably using the button.

Video files are named automatically using timestamps.

### Phase 3 – Slit-Lamp Integration (Optical and Mechanical)

Optical alignment and mechanical stability are inherently coupled in an observer-port imaging system. To meaningfully assess vignetting, focus, and field of view, the camera must be held in a repeatable physical relationship to the slit lamp observer eyepiece.

As a result, this phase treats optical alignment and mounting as a single integrated workstream, progressing from temporary, fully reversible fixtures toward a unified, mountable housing that contains all system components.

This phase is deliberately structured into sub-phases to allow early exit before committing to custom mechanical parts if optical feasibility is not validated.

### Phase 3a – Rapid Alignment and Temporary Fixtures

Objective: To establish a viable optical geometry between the camera sensor and the slit lamp observer eyepiece using temporary, reversible mounting methods.

Key Characteristics:

- No permanent modification to slit lamp

- No custom printed housing

- Low cost, rapid iteration

- Fully reversible

Tasks:

- Measure observer eyepiece outer diameter, length, and geometry

- Measure camera module stack height and lens protrusion

- Use temporary fixtures (e.g. rubber eyepiece sleeves, foam spacers, hose clamps)

- Adjust axial distance and alignment to minimise vignetting

- Assess sensitivity to tilt and axial displacement

Outputs:

- Target axial spacing from eyepiece exit pupil

- Acceptable alignment tolerance

- Baseline image quality reference

- Confirmation of optical feasibility using compact camera hardware

Acceptance Criteria: Usable image with minimal vignetting is achieved.

Image can be reproduced after removal and re-attachment using temporary fixtures

### Phase 3b: Design & Mechanical Development (Mount)

Objective: To translate the validated optical geometry into a repeatable mechanical mount using parametric design and rapid prototyping.

Key Characteristics:

- First custom mechanical components

- Alignment repeatability over aesthetics

- Adjustable features retained

Tasks:

- Create parametric CAD model (FreeCAD) for:

  - Eyepiece sleeve or collar

  - Camera carrier aligned to optical axis

- Incorporate axial adjustment or shimming

- Add anti-tilt and alignment-preserving features

- Produce iterative prototypes via 3D printing

- Test fit, stability, and repeatability on slit lamp

Acceptance Criteria:

- Mount maintains optical alignment across multiple attach/remove cycles

- No damage to slit lamp components

- Image quality matches or improves upon Phase 3A baseline

### Phase 3c – Integrated Housing Prototype

Objective: To design and prototype a single mountable housing that integrates all system components while preserving optical alignment and clinical usability.

Integrated Components:

- Camera module and lens

- Raspberry Pi Zero 2 WH

- Illuminated physical button

- USB storage access or internal USB device

- Camera ribbon cable and power routing

Design Principles:

- Optical axis defined by the eyepiece interface

- All electronics referenced mechanically to this axis

- Fully removable and non-invasive

- Serviceable access to SD card, USB, and connectors

- Adequate ventilation and thermal considerations

- Strain relief for all cables

Tasks:

- Extend parametric CAD model to include electronics enclosure

- Design internal mounting features for Pi, camera, and button

- Route cables to avoid strain or optical interference

- Produce integrated housing prototypes via 3D printing

- Validate full system operation while mounted

Acceptance Criteria:

- Complete housing mounts securely and repeatably

- One-button headless operation functions as intended

- Video recording is stable and reproducible

- System can be safely mounted and removed without tools

Notes on CAD Prototyping:

Parametric CAD modelling and rapid prototyping are used to enable fast iteration and adaptation to different slit lamp geometries. Mechanical design prioritises alignment stability and usability over cosmetic finish during this proof-of-concept phase.

### Phase 4 – File Handling and Data Flow

Objective: To enable simple retrieval and sharing of recorded videos.

File Strategy:

- Record video in H.264 format.

- Optionally convert recordings to MP4 automatically after capture.

Directory Structure:

- USB root directory contains a slit lamp recordings folder.

- Each recording is saved with a timestamp-based filename.

No patient-identifying data will be stored during the proof-of-concept phase.

### Phase 5 – Reliability and Workflow Enhancements

Objective: To improve robustness and suitability for demonstrations.

Potential Enhancements:

- Read-only operating system configuration.

- Watchdog service for automatic restart on failure.

- Visual indicator to show recording status.

- Controlled and safe shutdown procedure.

### Risks and Mitigations

- Optical misalignment risk is mitigated by adjustable mounting and iterative testing.

- Data loss risk is mitigated by controlled shutdown and robust file handling.

- Overheating risk is mitigated by passive ventilation and low-power hardware.

- User error risk is mitigated by a single-button interface.

### Estimated Timeline

Week 1: Hardware bring-up and basic testing.

Week 2: Button-controlled recording implementation.

Week 3: Optical alignment and initial mounting trials

Week 4: Integrated mount and housing prototyping

Week 5: Demonstration and review.

### Cost Estimate (Proof of Concept)

| Item | Cost |
| --- | --- |
| Raspberry Pi Zero 2WH | €25 |
| Camera Module | €7.50 - €30 |
| Cabling | €5 |
| Mounting | €2 |
| Buttons | €2 |
| Optional: 3D Printer | €210 |
| Optional: 3D Printing Service (External) | €50-200 |

Estimated total cost is €150 – 300.